## Slide 1: LESSONS LEARNT FROM THE GAMMA-BASED MONITORING STATIONS OF R.A.R.E. : DEVELOPMENT AND REAL-TIME PERFORMANCE
José Ángel Corbacho Merino
Environmental Radioactivity Laboratory. University of Extremadura
Cáceres, SPAIN
Workshop 2019, 3-5 April 2019, Roskilde, Denmark

## Slide 2
Environmental Radiological Network of Extremadura (R.A.R.E.) Located in Extremadura, SW SPAIN
  A brieft presentation
RADIOLOGICAL PARAMETERS
  Essential parameter: H*(10)  drawbacks
  RARE´S Gamma Spectrometry Systems
SOFTWARE APPLICATIONS
CONCLUSIONS
Outline
Notes: Justificar la conveniencia de los SGC en el ámbito de la redes de alerta

## Slide 3
RARE NETWORK
Commission date: September 1, 1983
Power generation:
Units operational: 1 × 1011 MW + 1 × 1006 MW
Capacity factor: 85.88%
Annual net output:15,174 GW·h (2016)
ALMARAZ NUCLEAR POWER PLANT (ANPP)
Notes: RARE AUTOMATIC RADIATION MONITORING NETWORK

## Slide 4
NATURAL BACKGROUND
mR/h

## Slide 5
- H*(10) monitors
- water monitoring stations
- airborne monitoring stations
STARTING POINT (1990)

## Slide 6
CURRENT STATE (2019)
2 monitoring drones
(under development)
1 monitoring
Mobile Laboratory
ISP/IEC 17025
UNE-EN ISO 9001
17 RARE H*(10)
monitoring stations
+
5 RARE air
monitoring stations
+
3 RARE water
monitoring stations

## Slide 7
Almaraz NPP uses for cooling the international Tagus River, that runs into Portugal.
Almaraz NPP vicinity

## Slide 8
The data are sent to RARE’s headquarters in quasi real-time, using at least two communication routes in Almaraz NPP area.
TRANSMISION MODALITIES
Digital radio
Analogical radio
Intranet
Twisted pair
Internet
Satellite  Mobile lab
COMMUNICATIONS: ESSENTIAL TOOL.

## Slide 9
RADIOLOGICAL PARAMETERS
Ambient equivalent dose rate: H*(10) :
fundamental parameter in radioprotection field:
COUNCIL DIRECTIVE 1996/29/EUROATOM
It is usually measured by a Geiger-Müller counter that is a sensitive (over an extended operational range of dose), robust, cheap and easy to use instrument. Therefore, dense networks are available.
It can be considered as a global parameter since most of early warning networks in Europe and world-wide comprise automatic measurement stations equipped with GM tubes.

## Slide 10
H10 (μSv/h)
.
CSN
EURDEP
RARE
Geiger-Müller counters
Notes: http://www.larescvalenciana.org/iso-9001-en-residencias-de-ancianos/

## Slide 11
RADIOLOGICAL PARAMETERS
OUR EXPERIENCE OVER THE LAST 30 YEARS USING G-M PROBES

## Slide 12
A GM tube can only detect the presence and intensity of radiation.
Geiger-Müller´s drawback
Increment ambient equivalent dose rate due to the scavenging of the airborne naturally occurring radionuclides during precipitation events is a well-known phenomenon
FACTS TO THINK ABOUT
Could a simultaneous weak man-made radiological anomaly overlook?
Dose rate (mSv/h)
Dose rate (mSv/h)
Precipitation (mm)
Notes: https://www.youtube.com/watch?v=WZ7Bo5CJCik
https://www.enac.es/plan-transicion-iso-17025
http://arrizabalagauriarte.com/iso-iec-170252017-caracteristicas-novedades-la-version-2017/
https://www.ambientalys.com/iso-iec-17025-2017

## Slide 13
2) Anomalous and continuous increment of H*(10) during summer
GM wrong behaviour due to high temperatures
The damaged monitor was replaced
Was it a real radiological alteration?
Dose rate increase due to higher Rn-222 concentration in air during summer

## Slide 14
The driver of a vehicle with radioactive material stopped for a while at the `petrol station….
3) Electronic noise or radiological anomaly?
Spurious increment
(electronic noise)
Mobile lab at a petrol station
Was it a real radiological alteration?

## Slide 15
SAUCEDILLA monitoring station during the days when Fukushima’s radioactive cloud reached Europe
3)Low sensitivity
The average daily value of the dose rate monitors did not detect alterations due to radioactive cloud
Amax(131I-part.)= 3 mBq/m3

## Slide 16
THE INFORMATION PROVIDE BY G-M PROBES MUST BE COMPLEMENTED BY OTHER MONITORING SYSTEM

## Slide 17
COMMERCIAL SYSTEM
Particulate fraction: gross a and b parameters (Rn222 compensation algorithm)
Gaseous fraction: I-131
FORMER RARE´S ATMOSPHERIC MONITORING STATION
Negative gross alpha and beta activity?

## Slide 18
Could man-made beta emitters be detected in the airborne fraction ?
Gross beta airborne activity
ANSWER: They were not detected because the natural gross beta oscillations due Rn-222 daughters were higher
WE COME BACK, AGAIN, TO FUKUSHIMA’S RADIOACTIVE CLOUD OVER RARE’S STATIONS
Higher activity detected

## Slide 19
FORMER RARE´S WATER MONITORING STATION
COMMERCIAL SYSTEM
Sludge accumulation at the bottom of the stainless steel vessel
Electronic system based on monochannel.
Continuous net I-131 and Cs-137 activity ?

## Slide 20
we opt for the gamma spectrometry systems
More compact electronics
High variety of gamma detectors

## Slide 21
RARE SYSTEM
Gamma spectrometry
COMMERCIAL SYSTEM
Particulate fraction:
a and b global parameters
Gaseous fraction:
131I
NaI(Tl) 2” x 2” or
LaBr3(Ce) 1.5” x 1.5”
NaI(Tl) 2” x 2”
RARE´S ATMOSPHERIC MONITORING STATION

## Slide 22
RARE SYSTEM
RARE´S ATMOSPHERIC MONITORING STATION

## Slide 23
RARE SYSTEM
COMMERCIAL SYSTEM
NaI(Tl) 2” x 2” or
LaBr3(Ce) 1.5” x 1.5”
RARE´S WATER MONITORING STATION
Gamma spectrometry

## Slide 24
RARE´S WATER MONITORING STATION
Rainfall scavenging

## Slide 25
RARE´S GAMMA SPECTROMETRY
routine operation
Early warning mode
(10min spectra)
and surveillance mode
(6h and 24h spectra)
Total gamma spectra at day:
141  10 min
4  6 h
1  24 h
 | MDA (Bq/m3) |  | 
 | 10 min | 6 h | 24 h
Ba-140 | 35 | 0,18 | 0,022
I-131 particulate | 3 | 0,018 | 0,004
Cs-137 | 5 | 0,025 | 0,003
Co-60 | 6 | 0,011 | 0,004
Zn-65 | 40 | 1,53 | 0,38
I-131 gaseous | 36 | 0,18 | 0,023
Airborne monitoring station
Flow: Particulate matter fraction (LaBr3): 25 m3/h
Gaseous fraction (NaI): 8 m3/h

## Slide 26
THE INFORMATION PROVIDED BY A RADIOLOGICAL NETWORK REQUIRES POWERFULL SOFTWARE
WE HAVE DEVELOPED OUR SOFTWARE TOOLS

## Slide 27
Vicinity of Almaraz NPP
Further Spanish locations
Portugal
Monitoring stations
Mobile lab
SERVER 1
SERVER 2
SERVER 3
Drone
Considering the number of monitoring stations and their analytical capabilities: 200000 data/day
(radiological, meteorological and operational)
SOFTWARE TOOLS
SOFTWARE APPLICATIONS
(backup)
(backup)

## Slide 28
H*(10)
Level 1: 30% of monthly average dose rate plus annual average dose rate from natural background.
Level 2: 0.114 μSv·h–1 plus annual average dose rate from natural background.
Level 3: 0.228 μSv·h–1 plus annual average dose rate from natural background.
Station state
Working
Shut down
LEVEL 1

## Slide 29
Activity concentration (green)
and
MDA
(orange)
Air flow and its proper operational range
Direct access to gamma spectrum for inspection
Activity concentrations

## Slide 30
Direction between the station and Almaraz NPP (pink lines)
Wind direction on real-time (blue points)
Meteorological data
WIND
RAIN AND
HUMIDITY
TEMPERATURE

## Slide 31
Flow, voltage,
indoor temperature
INPUT VOLTAGE
OUTPUT VOLTAGE
(stabilized)
Operational data
MONITORING OF UPS

## Slide 32
?
Easy identification of operational anomalies
Jammed filter
Flow decreases
Activities increase
Webcam image (visualized at RARE´s headquarters)

## Slide 33
Indoor temperature at FREGENAL has decreased. Now it is 29.4ºC
Message 22:49
Lack of power supply from the company
Air-conditioned shut down
Automatized electronic shut down (if T>30ºC)
This monitoring station doesn´t have a standby diesel power generator
Operational anomaly has finished
SMS on real time of radiological and operational anomalies
High indoor temperature at FREGENAL. During the last 30 min it is higher than 33ºC
Message 12:44
Shut down of the electronic at FREGENAL since 12:46:21
Message 13:47
T
H*(10)
VIN

## Slide 34
GAMMA SPECTROMETRY
OUR EXPERIENCE USING GAMMA SPECTROMETRY IN AUTOMATIC MONITORING STATION

## Slide 35
The main Gamma spectrometry drawbacks are:
1- Resolution of gamma scintillation detectors NaI(Tl) and LaBr3(Ce)
- The overlaping of photopeaks
I-131 (364 keV) -- Pb-214 (352 keV)
Cs-134 (604 keV) -- Bi-214 (609 keV)
Zn-65 (1115 keV) -- Bi-214 (1120 keV)
FALSE POSITIVES

## Slide 36
SOLUTION:
1- Stripping methods  Requires more complex calculation
2- Moved peak area method  Simple and easy calculation.
It solves I-131/Pb-214 and Cs-134/Bi-214 overlapping .
Determination of peak Area of Pb-214 (352 keV) from Pb-214 (295 keV)
Susbtraction of calculated peak area Pb-214 of overlapped peak Pb-214/I-131
Remove > 90% I-131 false positive activity

## Slide 37
2- Integration time vs sensitivity
We need an agreement between integration time and sensitivity
SOLUTION:
In early warning mode we have established a integration time of 600 s. But, if a radiological anomaly is detected our software allows to decrease this integration time remotely

## Slide 38
The only way to identify radioisotopes and quantify activity concentrations in quasi-real-time is to use automatic gamma spectrometry monitoring station.
RARE´s atmospheric monitoring stations are based not only on gamma dose rate determinations by Geiger-Müller counters (passive measurements) but also on sample retention and the gamma spectrometry of its gaseous and the particulate fractions.
All in situ gamma spectrometry system (passive or active devices) have got some drawbacks but it can be solved using powerful software tools.
Our gamma spectrometry has got the advantage that the geometry is well known and then we can get exact values for the activities.
The drawback of our system is the possibility that radioactive sources far to airborne collector can’t be detected.
Some and general conclusions

## Slide 39
Thank you very much
for your attention